Harmonics – Class-C per A-14(Run time)

EUT: Equipment under test	Tested by: Tested by

Test category: Class-C per A-14 (European limits)	Test Margin: 100

Test date: 2000-1-1	Start time: 0:02:04	End time: 0:03:14

Test duration (min): 1	Data file name: H-000179.cts_data

Comment: Comment

Customer: Customer information

Test Result: Pass	Source qualification: Normal

Current & voltage waveforms

Harmonics and Class C limit line (Table-3 Class-D)	European Limits

Test result:  Pass     or waveform fails phase angle spec. - see below

Current Test Result Summary (Run time)

EUT: Equipment under test	Tested by: Tested by

Test category: Class-C per A-14 (European limits)	Test Margin: 100

Test date: 2000-1-1	Start time: 0:02:04	End time: 0:03:14

Test duration (min): 1	Data file name: H-000179.cts_data

Comment:  Comment

Customer:  Customer information

Test Result: Pass	Source qualification:  Normal

THC(A): 0.01       I-THD(pk%): 9.86         POHC(A):  0.003        POHC Limit(A):  0.005

Highest parameter values during test:

V_RMS (Volts): 	120.10	Frequency(Hz):   60.00

I_Peak (Amps):	0.151	I_RMS (Amps):	0.093

I_Fund (Amps):	0.093	Crest Factor:	1.634

Power (Watts): 	11	Power Factor:	0.994

Harm#	Harms(avg)	100%Limit	%of Limit	Harms(max)	150%Limit	%of Limit	Status

2	0.001

3	0.006	0.038	16.1	0.006	0.057	11.42	Pass

4	0.000

5	0.004	0.021	16.8	0.004	0.032	12.42	Pass

6	0.000

7	0.002	0.011	20.9	0.003	0.017	15.03	Pass

8	0.000

9	0.002	0.006	31.1	0.002	0.008	22.83	Pass

10	0.000

11	0.001	0.004	35.5	0.002	0.006	27.36	Pass

12	0.000

13	0.001	0.003	34.8	0.001	0.005	28.34	Pass

14	0.000

15	0.001	0.003	36.2	0.001	0.004	30.66	Pass

16	0.000

17	0.001	0.003	37.7	0.001	0.004	32.77	Pass

18	0.000

19	0.001	0.002	40.7	0.001	0.003	35.31	Pass

20	0.000

21	0.001	0.002	44.4	0.001	0.003	37.31	Pass

22	0.000

23	0.001	0.002	48.1	0.001	0.003	40.45	Pass

24	0.000

25	0.001	0.002	52.3	0.001	0.003	43.42	Pass

26	0.000

27	0.001	0.002	55.6	0.001	0.002	46.78	Pass

28	0.000

29	0.001	0.001	59.1	0.001	0.002	49.49	Pass

30	0.000

31	0.001	0.001	63.7	0.001	0.002	54.87	Pass

32	0.000

33	0.001	0.001	68.0	0.001	0.002	57.23	Pass

34	0.000

35	0.001	0.001	71.3	0.001	0.002	59.48	Pass

36	0.000

37	0.001	0.001	74.4	0.001	0.002	63.07	Pass

38	0.000

39	0.001	0.001	77.8	0.001	0.002	67.22	Pass

40	0.000

Note: Dynamic limits were applied for this test (Class C < 25 Watts).

Voltage Source Verification Data (Run time)

EUT: Equipment under test	Tested by: Tested by

Test category: Class-C per A-14 (European limits)	Test Margin: 100

Test date: 2000-1-1	Start time: 0:02:04	End time: 0:03:14

Test duration (min): 1	Data file name: H-000179.cts_data

Comment:  Comment

Customer:  Customer information

Test Result: Pass	Source qualification: Normal

Highest parameter values during test:

Voltage (Vrms): 	120.10	Frequency(Hz):   60.00

I_Peak (Amps):	0.151	I_RMS (Amps):	0.093

I_Fund (Amps):	0.093	Crest Factor:	1.634

Power (Watts): 	11	Power Factor:	0.994

Harm#	Harmonics V-rms	Limit V-rms	% of Limit	Status

2	0.172	0.240	71.53	OK

3	0.310	1.081	28.72	OK

4	0.013	0.240	5.60	OK

5	0.055	0.480	11.47	OK

6	0.007	0.240	2.92	OK

7	0.015	0.360	4.06	OK

8	0.009	0.240	3.62	OK

9	0.016	0.240	6.52	OK

10	0.017	0.240	7.24	OK

11	0.010	0.120	8.37	OK

12	0.004	0.120	3.71	OK

13	0.007	0.120	5.80	OK

14	0.006	0.120	4.72	OK

15	0.008	0.120	6.75	OK

16	0.008	0.120	6.50	OK

17	0.005	0.120	3.94	OK

18	0.007	0.120	6.15	OK

19	0.003	0.120	2.34	OK

20	0.007	0.120	5.83	OK

21	0.004	0.120	3.59	OK

22	0.005	0.120	4.15	OK

23	0.003	0.120	2.25	OK

24	0.003	0.120	2.57	OK

25	0.005	0.120	3.85	OK

26	0.005	0.120	3.79	OK

27	0.003	0.120	2.79	OK

28	0.003	0.120	2.76	OK

29	0.004	0.120	3.15	OK

30	0.004	0.120	2.95	OK

31	0.003	0.120	2.73	OK

32	0.003	0.120	2.76	OK

33	0.005	0.120	4.17	OK

34	0.002	0.120	2.01	OK

35	0.003	0.120	2.27	OK

36	0.003	0.120	2.26	OK

37	0.003	0.120	2.64	OK

38	0.003	0.120	2.91	OK

39	0.004	0.120	3.65	OK

40	0.004	0.120	3.02	OK